NATIONAL INSTITUTE OF TECHNOLOGY,

TIRUCHIRAPPALLI - 620 015

Tender Notification No.9/2007

Department of Chemistry

Item No.08. Digital Polarimeter

Specifications for Automatic Digital Polarimeter

Light source: Sodium Lamp or Tungsten-Halogen lamp

Optical wavelength: 589nm

Measuring range: plus or minus 90 Deg.

Resolution: 0.001 Deg.

Repeatability: 0.002 Deg.

Accuracy: 0.01- 0.002°

Sample chamber: should accept sample tubes up to 200mm

Sample tubes: cylindrical glass cell

10mm ID x 20mm and 10mm ID x 50mm path length each one number.

Temperature range: 0-40° C (accuracy +/- 0.1 Deg C)

Readout modes: Optical rotation, optical specific rotation, concentration, optical purity

Display: on screen or on personal computer via USB ports (if operating on PC, PC requirement should be mentioned)

Calibration: Using Standards or In-built

Software: suitable software should be provided

Detector: PMT Tube

Aperture: should be variable for low concentration measurements

Response speed: should be +/- 6°/sec

The polarimeter should be up gradable on the field for other wavelength filters with appropriate light source.

NATIONAL INSTITUTE OF TECHNOLOGY,

TIRUCHIRAPPALLI - 620 015

Tender Notification No.9/2007

Department of Chemistry

Item No.09. BET Surface Area

Specification for BET Surface Area & Pore Size analyzer

Multi\ single point BET surface area

Measurement Range: 0.01- 2000 m2/ g

Measurement precision: Reproducibility ± 1%

Sample quantity: Sample capacity is approximately 2.0 cm3

Gas type: Pure nitrogen or mixed gas

Deaeration temperature: 450 Deg C

Calibration: Built –In calibration by injection

Analysis: Pore volume, Meso pore size distribution, Langmuir Surface area, Micro pore volume by T- plot method.4Analysis by both adsorption and desorption method.

Surface area reference material & standard accessories

PC & Printer for Control: Windows XP with suitable software.

The BET analyzer should be up gradable for further applications.

The accessories supplied along with the instruments like tubes, regulators, Gas cylinders, Liquid nitrogen containers etc., should be specified.